To find the best peak detection method and optimize it, three different approaches were utilized. These approaches were using data without filtering, with Savitzky-Golay filter, and with a Gaussian filter. Each non-normalized spectrum of each pixel was added together to create one global spectrum. The same thing happened to the TIC-normalized spectrum as well. Each approach was tested with two different inputs and combinations of parameters. After that, peak detection using the scipy library was performed. Then, peaks that were too close to a certain Dalton (Da) from the stronger peak were filtered out. Twice the size of the intended top m/z with the higher signal intensities was chosen. Then, the intensity of those top m/z was intensified by adding neighboring m/z intensities in a specific Dalton range. At the end, the top n m/zs with the highest signal intensities were chosen and saved as a txt file.

Each approach is going to be explained in detail accordingly.

- Without filtering:

The minimum peak height relative to the highest peak was set to 0.00001 for nomination for peak detection. The target was the top 1000 most intense peaks. The minimum distance between two consecutive chosen peaks was set to one of the numbers in the list [0.1, 0.05, 0.024, 0.02, 0.01], and the unit was Daltons. The maximum distance in Daltons to add neighboring m/zs’ intensities to the peak was chosen from the range [0.01, 0.02, 0.024, 0.03, 0.04, 0.05, 0.001, 0.0001, 0.00001, 0.000001].

- Gaussian filter:

The Gaussian filter from the scipy library was used in this approach. The standard deviation for the kernel was chosen from the range [0.2, 0.5, 1.0, 1.5, 2.0, 2.5, 3.0, 3.5, 4.0]. The threshold for peak nomination was set at 0.00001 times the height of the most intense peak. The goal was to identify the 1000 peaks with the highest intensities. The minimum spacing required between two selected peaks was chosen from the set [0.1, 0.05, 0.024, 0.02, 0.01] in units of Daltons. To aggregate neighboring m/z intensities into a detected peak, the maximum allowable distance in Daltons was selected from the list [0.0001, 0.001, 0.01, 0.02, 0.024, 0.03, 0.04, 0.05].

- Savitzky-Golay filter:

This approach utilized the Savitzky-Golay filter from the SciPy library, with the length of the filter window selected from the range [3, 5, 7, 9, 11, 13, 15, 17, 19, 21] and also the order of the polynomial used to fit the data was selected from the range [0, 1, 2, 3]. It is worth mentioning that the polynomial order should always be less than the window length. Peaks were nominated if their height reached at least 0.00001 of the maximum peak intensity. The objective was to extract the 1000 most intense peaks. The minimum required separation between consecutive selected peaks was drawn from the set [0.1, 0.05, 0.024, 0.02, 0.01] Daltons. To include nearby m/z values in a detected peak, the maximum distance for merging intensities was chosen from [0.01, 0.02, 0.024, 0.03, 0.04, 0.05, 0.001, 0.0001, 0.00001, 0.000001] Daltons.

These three approaches yielded 2365 .txt files. The results of the top 1000 most intense m/z produced by High Definition Imaging (HDI) software were used as a reference to find the best approach and its parameter combination. HDI uses its database to correct the reported m/zs from the experiment, so the reported m/z numbers differ slightly from the actual numbers. To compare the selected top m/zs with reference, a range of 0.02 Dalton difference was applied as an accepted tolerance threshold, and the unmatched m/zs were written in a separate .txt file for each combination. After comparison of the yielded 2365 files with the reference file, 2365 .txt files with unmatched m/zs were produced. Those files were sorted based on their volume. The smallest file showed the lowest number of unmatched m/zs. To put it better, it showed the highest matching m/zs. Based on the results, the peak detection method using Savitzky-Golay filter showed the best results by 86.5% matching rate. 12 parameter combinations for this filter showed the same matching results - 86.5%.

Input = non-normalized data

The maximum distance for merging intensities around peaks in Daltons = [0.001, 0.0001]

The minimum spacing between two selected peaks in Daltons = [0.024, 0.02, 0.01]

The polynomial order = [2, 3]

The length of the filter window = [9]

Overall, using the m/z values without intensity normalization as input resulted in a higher matching rate than using TIC-normalized values. As far as the maximum distance for merging intensities is concerned, the Mass Spectrometry instrument used for obtaining data had a minimum of 0.0043 Daltons data acquisition rate, so both 0.001 and 0.0001 meant that no addition happened, and these numbers were just safety measures. Since the minimum spacing between the selected peaks produced the same results, it was logical to choose the higher range, as it could aid in peak plotting, particularly in accounting for instrument drift. For polynomial order, 2 and 3 produced the same results; however, the order 3 polynomial made a smoother curve.

Based on the comparison results, the final version of the automatic peak detection code selects the top-N most intense m/z values without applying intensity normalization. In this version, the lowest nominated peak was approximately 10,000 times less intense than the highest, and a minimum spacing of 0.024 Daltons between consecutive peaks was enforced to avoid redundant detections. For signal smoothing, a Savitzky-Golay filter with a window length of 9 data points and a polynomial order of 3 was applied. Peak augmentation using neighboring points was not utilized during top-N selection. However, an optional feature was included to sum the intensities of n neighboring data points within a user-defined m/z window around each detected peak. This improvement enables the detection algorithm to capture and accommodate minor m/z drifts that may occur across different pixels during data acquisition.